## Slide 1: High performance liquid chromatography (HPLC)
Separate, identify & quantify components dissolved in a liquid solvent with a high analytical resolution. Sample carried by a moving media
  High-performance liquid chromatography (HPLC) is a separation technique that is used to analyze a wide variety of compounds.
  HPLC is based on the principle of partition chromatography, which is the separation of components of a mixture based on their different solubilities in two immiscible phases.
  In HPLC, the mobile phase is a liquid that flows through a column packed with a solid stationary phase.
  The components of the mixture are partitioned between the mobile and stationary phases, and the different components elute from the column at different times.

## Slide 2
High Performance Liquid Chromatography (HPLC) is a form of column chromatography that pumps a sample mixture or analyte in a solvent (known as the mobile phase) at high pressure through a column with chromatographic packing material (stationary phase).
The sample is carried by a moving carrier Media.
HPLC has the ability to separate, and identify compounds that are present in any sample that can be dissolved in a liquid in trace concentrations as low as parts per trillion.
Because of this versatility, HPLC is used in a variety of industrial and scientific applications, such as pharmaceutical, environmental, forensics, and chemicals.

## Slide 3: Retention Time:
The retention time of a solute is taken as the elapsed time between the time of injection of a solute and the time of elution of the maximum peak of that solute.
Sample retention time will vary depending on the interaction between the stationary phase, the molecules being analyzed, and the solvent, or solvents used.
As the sample passes through the column it interacts between the two phases at different rate, primarily due to different polarities in the analytes.
Analytes that have the least amount of interaction with the stationary phase or the most amount of interaction with the mobile phase will exit the column faster.

## Slide 5: Instrumentation:
Main components in an HPLC system include the solvent reservoir, or multiple reservoirs, a high-pressure pump, a column, injector system, the detector and waste.

## Slide 6
The reservoir holds the solvent, which is referred to as the mobile phase because it moves. There are usually a minimum of two reservoirs in a system, with each holding up to 1000 cc of solvent and usually fitted with a gas diffuser through which helium can be bubbled.
A pump is used to generate a specified flow of the mobile phase.
Although manual injection of samples is still possible, most HPLCs are now fully automated and controlled by computer.

## Slide 7: HPLC Pumps
Requirements of an Ideal HPLC Pump
While operating HPLC system there are certain requirements from pump like generation of pressure up to 5000 to 6000 psi without fluctuations. Pump should be suitable to a wide range of solvents providing a constant flow rate in the range of 0.1 – 10 ml per minute. Pumps should be made up of corrosion resistance material. Generally, the material of choice for the construction of the pump is Teflon 56, 57.
Type of Pumps used in HPLC
Reciprocating Pump
Screw Driven Syringe Pump or Displacement Pump
Pneumatic Pump or Constant Pressure Pump 35,56,57
Reciprocating Pump
It is the most prevalent and currently used pumping system on a commercial scale. It consists of a hydraulic chamber with a capacity varying from 35 µL to 400 µL. The chamber is connected to a motor-driven piston, which pumps the solvent in a back-and-forth motion. Check valves are present in a number of two, to regulate the solvent’s flow into and out of the column by alternatively opening and closing, respectively. A flexible diaphragm can also be used in a hydraulic fluid chamber to transmit the pumping action to the solvent. The use of a flexible diaphragm minimises the adulteration and erosion problems in the parts of the pump. In this, the movement of the piston causes the suction of the solvent from the mobile reservoir during the backward stroke and by the pressure of the pump, the solvent comes out of the column during the forward stroke and the inlet to the reservoir is closed. The main advantage of the reciprocating pump is that pressure up to 10000 psi can be applied, and some of the limitations are that it has a high maintenance cost and low discharging capacity

## Slide 8: Contd..
Displacement Pump or Syringe Pump (Screw Driven)
These pumps are attached to a digital stepping motor which works on the principle of a screw driven mechanism. The piston is triggered by a screw – feed drive through which solvent flows from the solvent chamber of finite capacity (250-500 ml). That is why the displacement pump is also known as the screw driven syringe type pump. This pump is capable of generating high pressure of 200-475 atm with the main advantage being that pulse loss flow is achieved but exhibits considerable inconvenience when the solvent composition is changed from isocratic to gradient flow 59,60.
Pneumatic Pump or Constant Pressure Pump
In a pneumatic pump, a gas cylinder is fitted, which creates the pressure through which the mobile phase is compelled through the column with the help of a large piston. The compression observed by the solvent is proportional to the ratio of the area of the two pistons, usually between 30:1 and 50:1, and a high pressure (1-400 atm) is generated with a low-pressure gas source of 1-10 atm. The arrangement of the valve permits the hasty refilling of the solvent chamber with a capacity of 70 ml. This system provides pulse loss and continuous pumping, having a high flow rate, so it is called a constant pressure pump

## Slide 9
The injector, or auto sampler, introduces the solvent into a phase stream that carries the sample into the high pressure (up to 400 bar) column, which contains specific packing material needed to effect separation. The packing material is referred to as the stationary phase because it is held in place by the column hardware.
A detector is needed to see the separated compound bands as they elute from the high pressure column.
The information is sent from the detector to a computer which generates the chromatogram.
The mobile phase exits the detector and is either sent to a waste, or collected, as desired.

## Slide 10: Detectors in HPLC
Bulk Property Detectors
Bulk property detectors are also known as universal detectors. In the mobile phase, it measures the variance in the physical properties of the solute related to the mobile phase alone. Generally, the applications of these types of detectors are found in abundance, but they have a limited range of detection and poor sensitivity. These detecting systems are easily exaggerated by small variations in the conformation of the mobile phase, preventing the use of techniques such as concentration gradient. For example, evaporative light scattering detector and refractive index detector 78,79.
Solute Property Detectors
Solute property detectors are also known as sensitive detectors. They are idyllically independent of the mobile phase, and counter to a specific chemical or physical characteristic of the solute. These generally show a wide linear response range and high sensitivity, but due to the fact that they are specific in nature and to meet the demands of an analytical column, more than one detector is required. For example, absorbance detector, fluorescence detector, electrochemical detector, and mass spectrometric detector 77.
Various types of detectors used in HPLC systems are described below:
UV/Visible Detector
Its principle behind detection is the absorption of light characteristics of the sample. It includes two types of detectors. One is a fixed wavelength detector and operates at a frequency of 254 nm, where most drugs comprising composites get absorbed, and the other one is a variable wavelength detector, operating at a frequency range of 190 to 600 nm. They provide good stability and are easy to operate. It detects or visualises the obtained results in two dimensions

## Slide 11: Contd…
Photodiode Array Detector
This is also used in UV spectrophotometer, and it works on the principle of light absorption; its operational frequency range falls between the wavelengths of 190 to 600nm. Simultaneously, radiation of all wavelengths falls on the detecting system. The analysis yields spectra with a three-dimensional plot of response vs. time vs. wavelength 82,83.
Fluorescence Detector
This detector works on fluorescent radiation emitted by compounds. Analyte atoms get excited by using a specific wavelength followed by the emission of fluorescence (which is considered a light signal). This detector has greater sensitivity and specificity, which is 10-1000 times greater than the UV detector. Even a single analyte fragment in the flow cell can be detected by using a fluorescence detector 84.
Conductivity Detector
It is based upon the electrical conductivity that is being consumed when the sample contains anions and captions as conducting ions. Therefore, detection is done on the basis of conducting ions in the sample and the response is recorded. Electronic resistance is measured and its value has a direct relationship with the number of ions present in the sample; if the presence of ions is high, they have a large value of electronic resistance and vice versa. Therefore, this can also be used for ion- exchange chromatography 39,79.
Amperometric Detector
When the potential is applied, it detects the compound based on its oxidation or reduction, followed by recording a diffusion current that is proportional to the concentration of the element eluted. This is related to the composites having functional groups that can be either oxidised or reduced. The sensitivity of this detector is very high 85.
Mass Spectrometry Detector
In this detector, a sample, analyte is detected on the basis of their molecular weight. Detecting system here is a mass analyser which detects the ions with respect to their mass to charge ratio, i.e., m/z ratio & amplifier, ions emerged from the mass analyser and record the responses. The analysis is useful for the identification of structural compounds. It can also be used for quantification of molecular and elemental compounds at very low detection limits

## Slide 12: Contd…
Refractive Index Detector
It is the only universal detector in HPLC. The detector works by measuring the variation in the refractive index of the column effluent passing through the flow cell. The refractive index difference between the mobile phase and the sample is directly proportional to the imbalance between them. Thus, the sensitivity is higher for a greater difference in refractive index. Due to its poor specificity and sensitivity, it is not much used in analytical applications .
Evaporative Light Scattering Detector
Less volatile compounds (semi-volatile compounds and non-volatile) in comparison to the mobile phase are easy to detect in an evaporative light scattering detector. It is mostly used for the analysis of compounds where UV detection must be restricted, or we can say, used for compounds which do not absorb UV radiation, such as antivirals, antibiotics

## Slide 13: Separation:
Injector
Detector
Column
Solvents
Mixer
Pumps
Chromatogram
Start Injection
mAU
time

## Slide 14: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 15: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 16: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 17: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 18: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 19: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 20: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 21: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 22: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 23: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 24: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 25: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 26: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 27: Separations
Injector
Detector
Column
Solvents
Pumps
Mixer
Chromatogram
Start Injection
mAU
time

## Slide 28: Types of HPLC:
Normal–phase chromatography.
Reversed-phase chromatography.

## Slide 29: Normal–phase chromatography
Normal–phase chromatography was one of the first kinds of HPLC that chemists developed. Also known as normal-phase HPLC (NP-HPLC) this method separates analytes based on their affinity for a polar stationary surface such as silica
NP-HPLC uses a non-polar, non-aqueous mobile phase (e.g. Chloroform), and works effectively for separating analytes readily soluble in non-polar solvents. The analyte associates with and is retained by the polar stationary phase

## Slide 30: Reversed-phase chromatography:
Reversed phase HPLC (RP-HPLC) has a non-polar stationary phase , moderately polar mobile phase. One common stationary phase is a silica which has been surface-modified with RMe2SiCl, where R is a straight chain alkyl group such as C18H37 or C8H17. With such stationary phases, retention time is longer for molecules which are less polar, while polar molecules elute more readily (early in the analysis). An investigator can increase retention times by adding more water to the mobile phase

## Slide 31: Stationary Phases:
Polar (Normal Phase): Silica, alumina
Non-Polar (Reversed Phase): ODS Silica gel, C18, C8

## Slide 32: Chromatography Stationary Phases
relatively polar surface
relatively nonpolar surface
Silica Gel
Derivatized Silica Gel
Where R = C18H37
hydrocarbon chain
(octadecylsilyl deriv.
silica or “C18”)
“normal phase”
“reversed phase”

## Slide 33: Normal vs. Reversed Phase Chromatography

## Slide 34: TLC vs. HPLC:

## Slide 35: Application of HPLC
1. Pharmaceuticals industry
To control the drug stability
Quantity of drug determination from pharmaceutical dosage forms, ex. Paracetamol determination in panadol tablet
Quantity of drug determination from biological fluids, ex: blood glucose level
2. Analysis of natural contamination
Phenol & Mercury from sea water
3. Forensic test
Determination of steroid in blood, urine & sweat.
4. Clinical test

## Slide 36: HPLC Applications
Chemical
Environmental
Pharmaceuticals
Consumer Products
Clinical
polystyrenes
dyes
phthalates
tetracyclines
corticosteroids
antidepressants
barbiturates
amino acids
vitamins
homocysteine
Bioscience
proteins
peptides
nucleotides
lipids
antioxidants
sugars
polyaromatic hydrocarbons
Inorganic ions
herbicides

## Slide 37: The Factors Which Influence The HPLC Performance:
Internal diameter of column
The smaller in diameter, the higher in sensitivity.
Pump pressure
The higher in pressure, the higher in separation.
Sample size
The polarity sample, solvent and column
Temperature
The higher in temperature, the higher in separation.

## Slide 38: Advantages:
Needs a small sample with a high accuracy and precis
Non-destructed sample during operation compared to GC.

## Slide 39: Disadvantages:
Need a skill to run the instruments.
Solvents consuming.

## Slide 40: Advances in HPLC
New Advances in High-Performance Liquid Chromatography Technique
Several advancements have been made in the classical HPLC technique and these newer techniques are Rapid Resolution Liquid Chromatography (RRLC), Ultra-Performance Liquid Chromatography (UPLC), Ultra-Fast Liquid Chromatography (UFLC) and Nano Liquid Chromatography (Nano LC). The description of these chromatographic techniques and their comparison with HPLC are summarized in Table 1 and discussed in the following sections